## The details
Report No.
Customer Name
Customer Address
| Model |  | (You can leave this field blank if you want to use the model from the tolerance Drop Down list)
10 | Tolerance
Customer ID Number
Serial Number
| Date Tested | 2025-01-28
| Date Due | 2026-01-28 | Interval | 1 | year
| Inspector | PE, ML
| Ambient Temperature | 75 | °F
| Ambient Humidity |  | %
| Test Location | Inspection
| On-Site? | Yes |  |  |  |  | TRUE
| Number of SMRs Cal'd: | 0 |  | Total pgs | 2 |  | FALSE
| Condition | Used
| Comments | This item was found in tolerance. No adjustments were performed.
| Result | Accepted
|  | Standard |  | Unistraint |  | Bar of Invar
| Uncertainty @ k=2: | 600 | µin | 600
| Standard Cal Date | 2024-06-05 |  | 2022-11-03
| Standard Due Date | 2025-06-05 |  | 2023-11-03
| Traceabiliy No. | 56561-059582 |  | 56561-035633
| CTE | 0.7 |  | 8.04 |  | 0.67
| Reference bar Length @ 68 degrees | 90.5402 | inches | 90.5413 |  | 90.5402
|  | 2299.7202 | mm | 2299.7485 |  | 2299.7202
A2LA ACCREDITED CERT
CERTIFICATE OF INSPECTION (NO "ACCEPTED" OR "REJECTED")

## JGI Method Testing
[14 empty] | INVAR bar | Flat | Clamped |  | CTE
| JGI METHOD |  |  |  | 10 | [8 empty] | 1 to 5 | 90.5482 | 90.5539 |  | 0.727
[6 empty] | 2299.7315 | Metric Length of standard, mm | [6 empty] | 1 to 4 | 67.9149 | 67.9191 |  |  | 1 | avg error | max error
Point to Point Volumetric Testing |  |  |  | 90.5406 | -90.5406 | 0 | [5 empty] | 75 |  | 1 to 3 | 45.2739 | 45.2757 |  | -90.5406 | -90.5406 | -90.5406 | -90.5406
| Artifact | Artifact Distance | Rotational Position | [10 empty] | 1 to 2 | 22.6372 | 22.6389 |  | 90.5406 | 90.5406
Test Position | Orientation | (Approximate) | of tracker head | Ref. Length | Greatest Dev. | Tolerance | distance (ft) | Result (P/F) | Result 1 | Result2 | Result3 | Temperature | CTE Correction applied |  | R.L. Value | R.L. CTE |  | Max | Min | Abs Max
1 | Horizontal | 2.7m / 9ft | 0° | 90.5406 | -90.5406 | 0.0025 | 9 | Fail |  |  |  | 75 | 0.00044365 | 7 | 90.5402 | 0.7 |  | -90.5406 | -90.5406 | -90.5406 | 2.7
2 | Horizontal | 2.7m / 9ft | 90° | 90.5406 | -90.5406 | 0.0025 | 9 | Fail |  |  |  | 75 | 0.00044365 | 7 | 90.5402 | 0.7 |  | -90.5406 | -90.5406 | -90.5406 | 2.7
3 | Horizontal | 2.7m / 9ft | 180° | 90.5406 | -90.5406 | 0.0025 | 9 | Fail |  |  |  | 75 | 0.00044365 | 7 | 90.5402 | 0.7 |  | -90.5406 | -90.5406 | -90.5406 | 2.7
4 | Horizontal | 2.7m / 9ft | 270° | 90.5406 | -90.5406 | 0.0025 | 9 | Fail |  |  |  | 75 | 0.00044365 | 7 | 90.5402 | 0.7 |  | -90.5406 | -90.5406 | -90.5406 | 2.7
5 | Vertical | 2.7m / 9ft | 0° | 90.5406 | -90.5406 | 0.0025 | 9 | Fail |  |  |  | 75 | 0.00044365 | 7 | 90.5402 | 0.7 |  | -90.5406 | -90.5406 | -90.5406 | 2.7
6 | Vertical | 2.7m / 9ft | 90° | 90.5406 | -90.5406 | 0.0025 | 9 | Fail |  |  |  | 75 | 0.00044365 | 7 | 90.5402 | 0.7 |  | -90.5406 | -90.5406 | -90.5406 | 2.7
7 | Vertical | 2.7m / 9ft | 180° | 90.5406 | -90.5406 | 0.0025 | 9 | Fail |  |  |  | 75 | 0.00044365 | 7 | 90.5402 | 0.7 |  | -90.5406 | -90.5406 | -90.5406 | 2.7
8 | Vertical | 2.7m / 9ft | 270° | 90.5406 | -90.5406 | 0.0025 | 9 | Fail |  |  |  | 75 | 0.00044365 | 7 | 90.5402 | 0.7 |  | -90.5406 | -90.5406 | -90.5406 | 2.7
9 | Right Diagonal 45° | 2.7m / 9ft | 0° | 90.5406 | -90.5406 | 0.0025 | 9 | Fail |  |  |  | 75 | 0.00044365 | 7 | 90.5402 | 0.7 |  | -90.5406 | -90.5406 | -90.5406 | 2.7
10 | Left Diagonal 45° | 2.7m / 9ft | 0° | 90.5406 | -90.5406 | 0.0025 | 9 | Fail |  |  |  | 75 | 0.00044365 | 7 | 90.5402 | 0.7 |  | -90.5406 | -90.5406 | -90.5406 | 2.7
11 | Right Diagonal 45° | 2.7m / 9ft | 90° | 90.5406 | -90.5406 | 0.0025 | 9 | Fail |  |  |  | 75 | 0.00044365 | 7 | 90.5402 | 0.7 |  | -90.5406 | -90.5406 | -90.5406 | 2.7
12 | Left Diagonal 45° | 2.7m / 9ft | 90° | 90.5406 | -90.5406 | 0.0025 | 9 | Fail |  |  |  | 75 | 0.00044365 | 7 | 90.5402 | 0.7 |  | -90.5406 | -90.5406 | -90.5406 | 2.7
13 | Compound 45 ° | 4.9m / 16ft | 0° | 90.5406 | -90.5406 | 0.0025 | 16 | Fail |  |  |  | 75 | 0.00044365 | 7 | 90.5402 | 0.7 |  | -90.5406 | -90.5406 | -90.5406 | 4.9
14 | Radial | 2.7m / 9ft | 0° | 90.5406 | -90.5406 | 0.0025 | 9 | Fail |  |  |  | 75 | 0.00044365 | 7 | 90.5402 | 0.7 |  | -90.5406 | -90.5406 | -90.5406 | 2.7
15 | Horizontal | 6.1m / 20ft | 0° | 90.5406 | -90.5406 | 0.0042 | 20 | Fail |  |  |  | 75 | 0.00044365 | 7 | 90.5402 | 0.7 |  | -90.5406 | -90.5406 | -90.5406 | 6.1
16 | Vertical | 6.1m / 20ft | 0° | 90.5406 | -90.5406 | 0.0042 | 20 | Fail |  |  |  | 75 | 0.00044365 | 7 | 90.5402 | 0.7 |  | -90.5406 | -90.5406 | -90.5406 | 6.1
17 | Horizontal | 12.2m / 40ft | 0° | 90.5406 | -90.5406 | 0.0075 | 40 | Fail |  |  |  | 75 | 0.00044365 | 7 | 90.5402 | 0.7 |  | -90.5406 | -90.5406 | -90.5406 | 12.2
18 | Vertical | 12.2m / 40ft | 0° | 90.5406 | -90.5406 | 0.0075 | 40 | Fail |  |  |  | 75 | 0.00044365 | 7 | 90.5402 | 0.7 |  | -90.5406 | -90.5406 | -90.5406 | 12.2
19 | Low/Backsighting | 2.7m / 9ft | 0° | N/A | 0 |  | 9 | [5 empty] | 0 |  | 90.5402 | 0.7 |  | #VALUE! | #VALUE! | #VALUE! | 2.7
20 | Mid/Backsighting | 2.7m / 9ft | 0° | N/A | 0 |  | 9 | [5 empty] | 0 |  | 90.5402 | 0.7 |  | #VALUE! | #VALUE! | #VALUE! | 2.7
21 | High/Backsighting | 2.7m / 9ft | 0° | N/A | 0 |  | 9 | [5 empty] | 0 |  | 90.5402 | 0.7 |  | #VALUE! | #VALUE! | #VALUE! | 2.7
22 | Mid/Backsighting | 1.8m / 6ft | 30° | N/A | 0 |  | 6 | [5 empty] | 0 |  | 90.5402 | 0.7 |  | #VALUE! | #VALUE! | #VALUE! | 1.8
23 | Mid/Backsighting | 1.8m / 6ft | 60° | N/A | 0 |  | 6 | [5 empty] | 0 |  | 90.5402 | 0.7 |  | #VALUE! | #VALUE! | #VALUE! | 1.8
24 | Mid/Backsighting | 1.8m / 6ft | 90° | N/A | 0 |  | 6 | [5 empty] | 0 |  | 90.5402 | 0.7 |  | #VALUE! | #VALUE! | #VALUE! | 1.8
25 | Mid/Backsighting | 2.7m / 9ft | 90° | N/A | 0 |  | 9 | [5 empty] | 0 |  | 90.5402 | 0.7 |  | #VALUE! | #VALUE! | #VALUE! | 2.7
26 | Mid/Backsighting | 6.1m / 20ft | 0° | N/A | 0 |  | 20 | [5 empty] | 0 |  | 90.5402 | 0.7 |  | #VALUE! | #VALUE! | #VALUE! | 6.1
PRELIMINARY READINGS (Prior to any adjustment, if the item is found out of tolerance) | [12 empty] | 0 |  | 90.5402 | 0.7 |  | 0 | 0 | 0
| Horizontal | 2.7m / 9ft | 0° | 90.5402 | -90.5402 | 0.0025 | 9 | Fail |  |  |  |  | 0 |  | 90.5402 | 0.7 |  | -90.5402 | -90.5402 | -90.5402 | 2.7 |  | 1 | avg error | max error
| Vertical | 2.7m / 9ft | 0° | 90.5402 | -90.5402 | 0.0025 | 9 | Fail |  |  |  |  | 0 |  | 90.5402 | 0.7 |  | -90.5402 | -90.5402 | -90.5402 | 2.7 | -90.5402 | -90.5402
| Right Diagonal 45° | 2.7m / 9ft | 0° | 90.5402 | -90.5402 | 0.0025 | 9 | Fail |  |  |  |  | 0 |  | 90.5402 | 0.7 |  | -90.5402 | -90.5402 | -90.5402 | 2.7 | 90.5402 | 90.5402
| Horizontal | 2.7m / 9ft | (select rotation) | 90.5402 | -90.5402 | 0.0025 | 9 | Fail |  |  |  |  | 0 |  | 90.5402 | 0.7 |  | -90.5402 | -90.5402 | -90.5402 | 2.7
Backsight Testing |  |  |  |  | Average Backsight Error | 0 | As Found Avg
|  | Artifact Distance | Rotational Position |  | Max Backsight error | 0 | As Found Max |  |  | Repetition 1 |  |  |  | Repetition 2 |  |  |  | Repitition 3
Test Position | Nest Position / Mode | (Approximate) | of tracker head | Avg X | Avg Y | Avg Z | Max O.A. |  | X | Y | Z |  | X | Y | Z |  | X | Y | Z
1 | Low / Frontsight | 1m / 3ft | 0° | #DIV/0! | #DIV/0! | #DIV/0! | As Left | As Found
2 | Low / Backsight | 1m / 3ft | 0° | #DIV/0! | #DIV/0! | #DIV/0! | 0
3 | Mid / Frontsight | 1m / 3ft | 0° | #DIV/0! | #DIV/0! | #DIV/0!
4 | Mid / Backsight | 1m / 3ft | 0° | #DIV/0! | #DIV/0! | #DIV/0! | 0
5 | High / Frontsight | 1m / 3ft | 0° | #DIV/0! | #DIV/0! | #DIV/0!
6 | High / Backsight | 1m / 3ft | 0° | #DIV/0! | #DIV/0! | #DIV/0! | 0
7 | Mid / Frontsight | 1m / 3ft | 30° | #DIV/0! | #DIV/0! | #DIV/0!
8 | Mid / Backsight | 1m / 3ft | 30° | #DIV/0! | #DIV/0! | #DIV/0! | 0
9 | Mid / Frontsight | 1m / 3ft | 60° | #DIV/0! | #DIV/0! | #DIV/0!
10 | Mid / Backsight | 1m / 3ft | 60° | #DIV/0! | #DIV/0! | #DIV/0! | 0
11 | Mid / Frontsight | 1m / 3ft | 90° | #DIV/0! | #DIV/0! | #DIV/0!
12 | Mid / Backsight | 1m / 3ft | 90° | #DIV/0! | #DIV/0! | #DIV/0! | 0
13 | Mid / Frontsight | 3m / 10ft | 90° | #DIV/0! | #DIV/0! | #DIV/0!
14 | Mid / Backsight | 3m / 10ft | 90° | #DIV/0! | #DIV/0! | #DIV/0! | 0
15 | Mid / Frontsight | 6m / 20ft | 0° | #DIV/0! | #DIV/0! | #DIV/0!
16 | Mid / Backsight | 6m / 20ft | 0° | #DIV/0! | #DIV/0! | #DIV/0! | 0
In-Line Distance
|  | Artifact Distance | Rotational Position | [5 empty] | ADM Results |  |  | Ref. Length | [8 empty] | Ref. Length
Test Position | Artifact Orientation/ Distance Function | (Approximate) | of tracker head | Ref. Length | ADM Result | Deviation | IFM Nom. | IFM Dev. | Result 1 | Result 2 | Result 3 | Temperature | CTE Correction applied |  |  | Ref. L. CTE |  | IFM 1 | IFM 2 | IFM 3 | Temperature | Avg IFM | CTE Correction applied
1 | Radial | 0.575m / 22in | 0° | 22.6369 | 22.6378 | 0.00089624 | 22.6367 | 0.0019743 | 22.6378 | 22.6378 | 22.6378 | 68.5 | 0.00013809 | 0.5 |  | 12.2 |  | 22.6385 | 22.6387 | 22.6389 | 71.1 | 22.6387 | 0.0008562 | 3.1
2 | Radial | 2.3m / 90 in | 0° | [9 empty] | 0 |  |  | 12.2 |  | 45.2768 | 45.277 | 45.277 | 70.3 | 45.2769 | 0.0012705 | 2.3
[13 empty] | 0 |  |  | 12.2 |  | 67.916 | 67.9162 | 67.9161 | 70 | 67.9161 | 0.0016572 | 2
|  |  |  | 90.554 | 90.5523 | -0.0016742 | 90.5535 | 0.0010515 | 90.5526 | 90.5523 | 90.5521 | 68.5 | 0.00055237 | 0.5 |  | 12.2 |  | 90.5545 | 90.5545 | 90.5545 | 70.6 | 90.5545 | 0.0028724 | 2.6
[13 empty] | 0 |  |  | 12.2 |  | 265.5803 | 265.5804 | 265.5803 | 71.2 | 265.5803 | 0.010368 | 3.2

## SMR Certification
Pg3 | Item | SN: | Position | X | Y | Z |  |  | Report Number: | 0 | [5 empty] | Page 3 of 2
| 1.5 inch SMR |  | Centered, Radial
| Tolerance: | 0.002 | Rotated Back | [8 empty] | Johnson Gage and Inspection, Inc.
| Stylus Probe? | yes | Rotated Forward | [8 empty] | 5920 West 21st Street North, Wichita, KS 67205
|  |  | Rotated Left | [8 empty] | Phone 316.943.7532 FAX 316.944.6256
Rotated Right
Laser Trackers- Reflector Certification Record
[10 empty] | Item: | 1.5 inch SMR |  |  | Date Tested: | 2025-01-28
[7 empty] | Result: | Pass |  | Serial Number: | 0 |  |  | Tested By: | PE, ML
|  |  | Max | 0 | 0 | 0 |  |  | Test | Reflector Position |  |  | X | Y | Z
|  |  | Min | 0 | 0 | 0 |  |  | 1 | Centered, Radial
|  |  | Max-Min/2 | 0 | 0 | 0 |  |  | 2 | Rotated Back
[9 empty] | 3 | Rotated Forward
[9 empty] | 4 | Rotated Left
Pg4 | [8 empty] | 5 | Rotated Right
| Item | SN: | Position | X | Y | Z
| 0.5 inch SMR |  | Centered, Radial
| Tolerance: | 0.0005 | Mirror rotated upward | [9 empty] | 0 | 0 | 0
| Stylus Probe? |  | Mirror rotated downard | [8 empty] | Minimum: | 0 | 0 | 0
|  |  | Rotated CW Horizontally | [8 empty] | Maximum: | 0 | 0 | 0
|  |  | Rotated CCW Horizontally | [8 empty] | (Max-Min)/2: | 0 | 0 | 0
Rotated 90° on beam axis
Rotated 180° on beam axis
Rotated 270° on beam axis
[7 empty] | Result: | Pass
|  |  | Max | 0 | 0 | 0
|  |  | Min | 0 | 0 | 0
|  |  | Max-Min/2 | 0 | 0 | 0
Pg5
[12 empty] | Overall Error (Max-Min)/2: | 0
| Item | SN: | Position | X | Y | Z | [5 empty] | Tolerance: | 0.002
| Leica T-Probe II |  | Centered, Radial | [8 empty] | Result: | Pass
| Tolerance: | 0.002 | Rotated Back
| Stylus Probe? | Yes | Rotated Forward | [5 empty] | Results displayed are in inches
|  |  | Rotated Left | [5 empty] | Report Number: | 0 | [5 empty] | Page 4 of 2
Rotated Right
Johnson Gage and Inspection, Inc.
5920 West 21st Street North, Wichita, KS 67205
Phone 316.943.7532 FAX 316.944.6256
[7 empty] | Result: | Pass
|  |  | Max | 0 | 0 | 0 | [5 empty] | Laser Trackers- Reflector Certification Record
|  |  | Min | 0 | 0 | 0
|  |  | Max-Min/2 | 0 | 0 | 0 |  |  |  | Item: | 0.5 inch SMR |  |  | Date Tested: | 2025-01-28
[10 empty] | Serial Number: | 0 |  |  | Tested By: | PE, ML
[9 empty] | Test | Reflector Position |  |  | X | Y | Z
Pg6 | [8 empty] | 1 | Centered, Radial
[9 empty] | 2 | Mirror rotated upward
| Item | SN: | Position | X | Y | Z |  |  | 3 | Mirror rotated downard
| Leica T-Probe II |  | Centered, Radial | [5 empty] | 4 | Rotated CW Horizontally
| Tolerance: | 0.002 | Rotated Back | [5 empty] | 5 | Rotated CCW Horizontally
| Stylus Probe? | yes | Rotated Forward | [5 empty] | 6 | Rotated 90° on beam axis
|  |  | Rotated Left | [5 empty] | 7 | Rotated 180° on beam axis
|  |  | Rotated Right | [5 empty] | 8 | Rotated 270° on beam axis
[13 empty] | 0 | 0 | 0
[12 empty] | Minimum: | 0 | 0 | 0
|  |  |  | 0.0007 | 0.0004 | 0.0003 | [5 empty] | Maximum: | 0 | 0 | 0
[7 empty] | Result: | Pass |  |  |  | (Max-Min)/2: | 0 | 0 | 0
|  |  | Max | 0.0007 | 0.0004 | 0.0003
|  |  | Min | 0.0007 | 0.0004 | 0.0003 |  |  |  | Diameter
|  |  | Max-Min/2 | 0 | 0 | 0 |  |  | 1 | D1 |  |  | 0
[9 empty] | 2 | D2 |  |  | 0
[9 empty] | 3 | D3 |  |  | 0
Pg7 | [8 empty] | 4 | D4 |  |  | 0
[9 empty] | 5 | D5 |  |  | 0
| Item | SN: | Position | X | Y | Z
Centered, Radial
| Tolerance: |  | Mirror rotated upward | [8 empty] | Variation in Diameter: | 0
| Stylus Probe? | no | Mirror rotated downard | [8 empty] | Overall Error (Max-Min)/2: | 0
|  |  | Rotated CW Horizontally | [8 empty] | Tolerance: | 0.0005
D1 |  |  | Rotated CCW Horizontally | [8 empty] | Result: | Pass
D2 |  |  | Rotated 90° on beam axis
D3 |  |  | Rotated 180° on beam axis | [5 empty] | Results displayed are in inches
D4 |  |  | Rotated 270° on beam axis | [5 empty] | Report Number: | 0 | [5 empty] | Page 5 of 2
D5 | [6 empty] | Result: | Pass
Variation in Diameter: |  | 0 | Max | 0 | 0 | 0 | [5 empty] | Johnson Gage and Inspection, Inc.
|  |  | Min | 0 | 0 | 0 | [5 empty] | 5920 West 21st Street North, Wichita, KS 67205
|  |  | Max-Min/2 | 0 | 0 | 0 | [5 empty] | Phone 316.943.7532 FAX 316.944.6256
Laser Trackers- Reflector Certification Record
Pg8
[10 empty] | Item: | Leica T-Probe II |  |  | Date Tested: | 2025-01-28
| Item | SN: | Position | X | Y | Z |  |  |  | Serial Number: | 0 |  |  | Tested By: | PE, ML
|  |  | Centered, Radial | [5 empty] | Test | Reflector Position |  |  | X | Y | Z
| Tolerance: |  | Mirror rotated upward | [5 empty] | 1 | Centered, Radial
| Stylus Probe? | no | Mirror rotated downard | [5 empty] | 2 | Rotated Back
|  |  | Rotated CW Horizontally | [5 empty] | 3 | Rotated Forward | [6 empty] | tprobe | smr
D1 |  |  | Rotated CCW Horizontally | [5 empty] | 4 | Rotated Left | [6 empty] | Centered, Radial | Centered, Radial
D2 |  |  | Rotated 90° on beam axis | [5 empty] | 5 | Rotated Right | [6 empty] | Rotated Back | Mirror rotated upward
D3 |  |  | Rotated 180° on beam axis | [13 empty] | Rotated Forward | Mirror rotated downard
D4 |  |  | Rotated 270° on beam axis | [13 empty] | Rotated Left | Rotated CW Horizontally
D5 | [6 empty] | Result: | Pass | [8 empty] | Rotated Right | Rotated CCW Horizontally
Variation in Diameter: |  | 0 | Max | 0 | 0 | 0 | [6 empty] | 0 | 0 | 0 |  |  | Rotated 90° on beam axis
|  |  | Min | 0 | 0 | 0 | [5 empty] | Minimum: | 0 | 0 | 0 |  |  | Rotated 180° on beam axis
|  |  | Max-Min/2 | 0 | 0 | 0 | [5 empty] | Maximum: | 0 | 0 | 0 |  |  | Rotated 270° on beam axis
[12 empty] | (Max-Min)/2: | 0 | 0 | 0
Diameter
D1
D2
D3
D4
D5
Variation in Diameter:
[12 empty] | Overall Error (Max-Min)/2: | 0
[12 empty] | Tolerance: | 0.002
[12 empty] | Result: | Pass
Results displayed are in inches
[9 empty] | Report Number: | 0 | [5 empty] | Page 6 of 2
Johnson Gage and Inspection, Inc.
5920 West 21st Street North, Wichita, KS 67205
Phone 316.943.7532 FAX 316.944.6256
Laser Trackers- Reflector Certification Record
[10 empty] | Item: | Leica T-Probe II |  |  | Date Tested: | 2025-01-28
[10 empty] | Serial Number: | 0 |  |  | Tested By: | PE, ML
[9 empty] | Test | Reflector Position |  |  | X | Y | Z
[9 empty] | 1 | Centered, Radial
[9 empty] | 2 | Rotated Back
[9 empty] | 3 | Rotated Forward
[9 empty] | 4 | Rotated Left
[9 empty] | 5 | Rotated Right
[13 empty] | 0.0007 | 0.0004 | 0.0003
[13 empty] | 0 | 0 | 0
[12 empty] | Minimum: | 0.0007 | 0.0004 | 0.0003
[12 empty] | Maximum: | 0.0007 | 0.0004 | 0.0003
[12 empty] | (Max-Min)/2: | 0 | 0 | 0
[12 empty] | Overall Error (Max-Min)/2: | 0
[12 empty] | Tolerance: | 0.002
[12 empty] | Result: | Pass
Results displayed are in inches
[9 empty] | Report Number: | 0 | [5 empty] | Page 7 of 2
Johnson Gage and Inspection, Inc.
5920 West 21st Street North, Wichita, KS 67205
Phone 316.943.7532 FAX 316.944.6256
Laser Trackers- Reflector Certification Record
[10 empty] | Item: | 0 |  |  | Date Tested: | 2025-01-28
[10 empty] | Serial Number: | 0 |  |  | Tested By: | PE, ML
[9 empty] | Test | Reflector Position |  |  | X | Y | Z
[9 empty] | 1 | Centered, Radial
[9 empty] | 2 | Mirror rotated upward
[9 empty] | 3 | Mirror rotated downard
[9 empty] | 4 | Rotated CW Horizontally
[9 empty] | 5 | Rotated CCW Horizontally
[9 empty] | 6 | Rotated 90° on beam axis
[9 empty] | 7 | Rotated 180° on beam axis
[9 empty] | 8 | Rotated 270° on beam axis
[13 empty] | 0 | 0 | 0
[12 empty] | Minimum: | 0 | 0 | 0
[12 empty] | Maximum: | 0 | 0 | 0
[12 empty] | (Max-Min)/2: | 0 | 0 | 0
Diameter
[9 empty] | 1 | D1 |  |  | 0
[9 empty] | 2 | D2 |  |  | 0
[9 empty] | 3 | D3 |  |  | 0
[9 empty] | 4 | D4 |  |  | 0
[9 empty] | 5 | D5 |  |  | 0
[12 empty] | Variation in Diameter: | 0
[12 empty] | Overall Error (Max-Min)/2: | 0
[12 empty] | Tolerance: | 0
[12 empty] | Result: | Pass
Results displayed are in inches
[9 empty] | Report Number: | 0 | [5 empty] | Page 8 of 2
Johnson Gage and Inspection, Inc.
5920 West 21st Street North, Wichita, KS 67205
Phone 316.943.7532 FAX 316.944.6256
Laser Trackers- Reflector Certification Record
[10 empty] | Item: | 0 |  |  | Date Tested: | 2025-01-28
[10 empty] | Serial Number: | 0 |  |  | Tested By: | PE, ML
[9 empty] | Test | Reflector Position |  |  | X | Y | Z
[9 empty] | 1 | Centered, Radial
[9 empty] | 2 | Mirror rotated upward
[9 empty] | 3 | Mirror rotated downard
[9 empty] | 4 | Rotated CW Horizontally
[9 empty] | 5 | Rotated CCW Horizontally
[9 empty] | 6 | Rotated 90° on beam axis
[9 empty] | 7 | Rotated 180° on beam axis
[9 empty] | 8 | Rotated 270° on beam axis
[13 empty] | 0 | 0 | 0
[12 empty] | Minimum: | 0 | 0 | 0
[12 empty] | Maximum: | 0 | 0 | 0
[12 empty] | (Max-Min)/2: | 0 | 0 | 0
Diameter
[9 empty] | 1 | D1 |  |  | 0
[9 empty] | 2 | D2 |  |  | 0
[9 empty] | 3 | D3 |  |  | 0
[9 empty] | 4 | D4 |  |  | 0
[9 empty] | 5 | D5 |  |  | 0
[12 empty] | Variation in Diameter: | 0
[12 empty] | Overall Error (Max-Min)/2: | 0
[12 empty] | Tolerance: | 0
[12 empty] | Result: | Pass
Results displayed are in inches
[9 empty] | Report Number: | 0 | [5 empty] | Page 3 of 2
Johnson Gage and Inspection, Inc.
5920 West 21st Street North, Wichita, KS 67205
Phone 316.943.7532 FAX 316.944.6256
Laser Trackers- Reflector Certification Record
[10 empty] | Item: | 0 |  |  | Date Tested: | 2025-01-28
[10 empty] | Serial Number: | 0 |  |  | Tested By: | PE, ML
[9 empty] | Test | Reflector Position |  |  | X | Y | Z
[9 empty] | 1 | Centered, Radial
[9 empty] | 2 | 0
[9 empty] | 3 | 0
[9 empty] | 4 | 0
[9 empty] | 5 | 0
[9 empty] | 6 | 0
[9 empty] | 7 | 0
[9 empty] | 8 | 0
[13 empty] | 0 | 0 | 0
[12 empty] | Minimum: | 0 | 0 | 0
[12 empty] | Maximum: | 0 | 0 | 0
[12 empty] | (Max-Min)/2: | 0 | 0 | 0
0
0
0
0
0
0
[12 empty] | Overall Error (Max-Min)/2: | 0
[12 empty] | Tolerance: | 0
[12 empty] | Result: | Pass
Results displayed are in inches
[9 empty] | Report Number: | 0 | [5 empty] | Page 3 of 2
Johnson Gage and Inspection, Inc.
5920 West 21st Street North, Wichita, KS 67205
Phone 316.943.7532 FAX 316.944.6256
Laser Trackers- Reflector Certification Record
[10 empty] | Item: | 0 |  |  | Date Tested: | 2025-01-28
[10 empty] | Serial Number: | 0 |  |  | Tested By: | PE, ML
[9 empty] | Test | Reflector Position |  |  | X | Y | Z
[9 empty] | 1 | Centered, Radial
[9 empty] | 2 | 0
[9 empty] | 3 | 0
[9 empty] | 4 | 0
[9 empty] | 5 | 0
[9 empty] | 6 | 0
[9 empty] | 7 | 0
[9 empty] | 8 | 0
[13 empty] | 0 | 0 | 0
[12 empty] | Minimum: | 0 | 0 | 0
[12 empty] | Maximum: | 0 | 0 | 0
[12 empty] | (Max-Min)/2: | 0 | 0 | 0
0
0
0
0
0
0
[12 empty] | Overall Error (Max-Min)/2: | 0
[12 empty] | Tolerance: | 0
[12 empty] | Result: | Pass
Results displayed are in inches

## Cert Page
Johnson Gage and Inspection, Inc.
5920 West 21st Street North, Wichita, KS 67205
Phone: 316-943-7532; FAX: 316-944-6256
Certificate of Calibration
| Report Number: | 0 | [8 empty] | Page 1 of 2
| Customer Name: | 0 | [8 empty] | Form 2554-4
| Customer Address: | 0 | [8 empty] | Original Rev. 4-21-15
| Item: | Laser Tracker | [6 empty] | Inspector: | PE, ML
| Customer ID Number | 0 | [6 empty] | Date Calibrated: | 2025-01-28
| Model Number: | Leica AT930 | [6 empty] | Recalibration Date: | 2026-01-28
| Serial Number: | 0 | [6 empty] | Ambient Temperature: | 75 |  |  |  |  | FALSE
| Location: | Inspection |  |  | -Tested On-Site |  |  |  | Relative Humidity: | 0 |  |  |  |  | TRUE
...
|  | Controlling Specification(s): | ANSI/NCSLI Z540-1, ISO/IEC 17025
|  | Calibration Procedure(s): | WP251420
...
This item was tested using traceable reference lengths throughout its working volume. These tests were designed to check the vertical and horizontal angle encoders at varying amounts of travel, as well as the distance encoder(s). Backsight testing was performed throughout the measuring envelope. A summary of the results has been provided below.
| Test Type |  | Maximum Observed Error (at 9ft) |  |  |  |  | Average Observed Error
| Volumetric Testing: |  |  | -90.5406 |  |  |  | -90.5406
| Backsight Testing: |  |  | 0 |  |  |  | 0 | [8 empty] | FALSE | 0
| Radial Distance Testing: |  |  | -90.5406 |  |  |  | -90.5406
[17 empty] | FALSE | 0
[17 empty] | FALSE | 0
[17 empty] | FALSE | 0
[17 empty] | FALSE | 0
Uncertainty @ k=2: 600 µin
| Detailed measurement results are included on page 2 | [15 empty] | The reported expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k such that the coverage probability corresponds to approximately 95%. This estimate was performed in accordance with guidelines set forth in ANSI/NCSL Z540-2. The acceptance or rejection of the item(s) is based on the actual test values shown, without adjustment for measurement uncertainty. All testing performed using standards traceable to NIST or to intrinsic standards. This document shall not be reproduced except in full, without the written permission of Johnson Gage and Inspection. Tolerance source key: C = Customer; S = Specification; M = Manufacturer; R = JGI Recommended.
The reported expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k such that the coverage probability corresponds to approximately 95%. This estimate was performed in accordance with guidelines set forth in ANSI/NCSL Z540-2. The acceptance or rejection of the item(s) is based on the actual test values shown, without adjustment for measurement uncertainty. All testing performed using standards traceable to NIST or to intrinsic standards. Tests Marked with an Asterisk are unaccredited. This document shall not be reproduced except in full, without the written permission of Johnson Gage and Inspection. Tolerance source key: C = Customer; S = Specification; M = Manufacturer; R = JGI Recommended.
The reported expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k such that the coverage probability corresponds to approximately 95%. This estimate was performed in accordance with guidelines set forth in ANSI/NCSL Z540-2. The acceptance or rejection of the item(s) is based on the actual test values shown, without adjustment for measurement uncertainty. All testing performed using standards traceable to NIST or to intrinsic standards. This document shall not be reproduced except in full, without the written permission of Johnson Gage and Inspection. Tolerance source key: C = Customer; S = Specification; M = Manufacturer; R = JGI Recommended.
...
Comments: | This item was found in tolerance. No adjustments were performed.
[16 empty] | Accredited Statement |  | The reported expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k such that the coverage probability corresponds to approximately 95%. This estimate was performed in accordance with guidelines set forth in ANSI/NCSL Z540-2. The acceptance or rejection of the item(s) is based on the actual test values shown, without adjustment for measurement uncertainty. All testing performed using standards traceable to NIST or to intrinsic standards. This document shall not be reproduced except in full, without the written permission of Johnson Gage and Inspection. Tolerance source key: C = Customer; S = Specification; M = Manufacturer; R = JGI Recommended.
[16 empty] | Unacc. Statement: |  | All testing performed using standards traceable to NIST or to intrinsic standards. This document shall not be reproduced except in full, without the written permission of Johnson Gage and Inspection.
Condition: | Used |  |  |  | Result: | Accepted | [9 empty] | 1
CALIBRATION STANDARDS | [12 empty] | Cal Date | Due: |  |  | Item | Serial | TN |  | Concat
| Standard: Reference Length Bar, Serial No.: LTRB1 Certified: 06/05/2024 Due:06/05/2025 Traceability: 56561-059582 | [10 empty] | Reference Length Bar SN: LTRB1 | 2024-06-05 | 2025-06-05 |  | 1 | Reference Length Bar | LTRB1 | 56561-059582 | TRUE | Standard: Reference Length Bar, Serial No.: LTRB1 Certified: 06/05/2024 Due:06/05/2025 Traceability: 56561-059582 | 06/05/2024 | 06/05/2025
[12 empty] | Reference Length Bar SN: LTRB2 | 2016-12-11 | 2017-12-11 |  |  | Reference Length Bar | LTRB2 | 209919 | FALSE | Standard: Reference Length Bar, Serial No.: LTRB2 Certified: 12/11/2016 Due:12/11/2017 Traceability: 209919 | 12/11/2016 | 12/11/2017
[12 empty] | SN: | [7 empty] | FALSE | Standard: , Serial No.: Certified: 01/00/1900 Due:01/00/1900 Traceability: | 01/00/1900 | 01/00/1900
[12 empty] | SN: | [7 empty] | FALSE | Standard: , Serial No.: Certified: 01/00/1900 Due:01/00/1900 Traceability: | 01/00/1900 | 01/00/1900
[12 empty] | SN: | [7 empty] | FALSE | Standard: , Serial No.: Certified: 01/00/1900 Due:01/00/1900 Traceability: | 01/00/1900 | 01/00/1900
[12 empty] | SN: | [7 empty] | FALSE | Standard: , Serial No.: Certified: 01/00/1900 Due:01/00/1900 Traceability: | 01/00/1900 | 01/00/1900
[12 empty] | SN: | [7 empty] | FALSE | Standard: , Serial No.: Certified: 01/00/1900 Due:01/00/1900 Traceability: | 01/00/1900 | 01/00/1900
[12 empty] | SN: | [7 empty] | FALSE | Standard: , Serial No.: Certified: 01/00/1900 Due:01/00/1900 Traceability: | 01/00/1900 | 01/00/1900
[12 empty] | SN: | [7 empty] | FALSE | Standard: , Serial No.: Certified: 01/00/1900 Due:01/00/1900 Traceability: | 01/00/1900 | 01/00/1900
[12 empty] | SN: | [7 empty] | FALSE | Standard: , Serial No.: Certified: 01/00/1900 Due:01/00/1900 Traceability: | 01/00/1900 | 01/00/1900
********************************************************************************************************************************** | [11 empty] | SN: | [7 empty] | FALSE | Standard: , Serial No.: Certified: 01/00/1900 Due:01/00/1900 Traceability: | 01/00/1900 | 01/00/1900
[7 empty] | Quality Approval: |  |  |  |  | SN: | [7 empty] | FALSE | Standard: , Serial No.: Certified: 01/00/1900 Due:01/00/1900 Traceability: | 01/00/1900 | 01/00/1900

## Datasheet, JGI Method
Report Number: | 0 | [6 empty] | Page 2 of 2
|  |  |  | Tracker Serial Number: | 0 |  |  |  | #DIV/0!
Point to Point Volumetric Testing |  |  |  | The deviation at each position is calculated as the maximum of three readings
Test | Artifact | Artifact Distance | Rotational Position |  | Artifact
Position | Orientation | (Approximate) | of tracker head | Reference length | Temperature | Deviation | Tolerance | Result (P/F)
1 | Horizontal | 2.7m / 9ft | 0° | 90.5406 | 75 | -90.5406 | 0.0025 | Fail | 90.5406
2 | Horizontal | 2.7m / 9ft | 90° | 90.5406 | 75 | -90.5406 | 0.0025 | Fail | 90.5406
3 | Horizontal | 2.7m / 9ft | 180° | 90.5406 | 75 | -90.5406 | 0.0025 | Fail | 90.5406
4 | Horizontal | 2.7m / 9ft | 270° | 90.5406 | 75 | -90.5406 | 0.0025 | Fail | 90.5406
5 | Vertical | 2.7m / 9ft | 0° | 90.5406 | 75 | -90.5406 | 0.0025 | Fail | 90.5406
6 | Vertical | 2.7m / 9ft | 90° | 90.5406 | 75 | -90.5406 | 0.0025 | Fail | 90.5406
7 | Vertical | 2.7m / 9ft | 180° | 90.5406 | 75 | -90.5406 | 0.0025 | Fail | 90.5406
8 | Vertical | 2.7m / 9ft | 270° | 90.5406 | 75 | -90.5406 | 0.0025 | Fail | 90.5406
9 | Right Diagonal 45° | 2.7m / 9ft | 0° | 90.5406 | 75 | -90.5406 | 0.0025 | Fail | 90.5406
10 | Left Diagonal 45° | 2.7m / 9ft | 0° | 90.5406 | 75 | -90.5406 | 0.0025 | Fail | 90.5406
11 | Right Diagonal 45° | 2.7m / 9ft | 90° | 90.5406 | 75 | -90.5406 | 0.0025 | Fail | 90.5406
12 | Left Diagonal 45° | 2.7m / 9ft | 90° | 90.5406 | 75 | -90.5406 | 0.0025 | Fail | 90.5406
13 | Compound 45 ° | 4.9m / 16ft | 0° | 90.5406 | 75 | -90.5406 | 0.0025 | Fail | 90.5406
14 | Radial | 2.7m / 9ft | 0° | 90.5406 | 75 | -90.5406 | 0.0025 | Fail | 90.5406
15 | Horizontal | 6.1m / 20ft | 0° | 90.5406 | 75 | -90.5406 | 0.0042 | Fail | 90.5406
16 | Vertical | 6.1m / 20ft | 0° | 90.5406 | 75 | -90.5406 | 0.0042 | Fail | 90.5406
17 | Horizontal | 12.2m / 40ft | 0° | 90.5406 | 75 | -90.5406 | 0.0075 | Fail | 90.5406
18 | Vertical | 12.2m / 40ft | 0° | 90.5406 | 75 | -90.5406 | 0.0075 | Fail
Backsight Testing
Test | Artifact Orientation/ | Artifact Distance | Rotational Position | Point Separation Error
Position | Test Mode | (Approximate) | of tracker head
1 | Low | 2.7m / 9ft | 0° | 0 |  |  |  |  | 0
2 | Mid | 2.7m / 9ft | 0° | 0
3 | High | 2.7m / 9ft | 0° | 0
4 | Mid | 1.8m / 6ft | 30° | 0
5 | Mid | 1.8m / 6ft | 60° | 0
6 | Mid | 1.8m / 6ft | 90° | 0
7 | Mid | 2.7m / 9ft | 90° | 0
8 | Mid | 6.1m / 20ft | 0° | 0

## Log
Following up on Failed Trackers
AIT- Called and emailed. No traction yet.
D-J -- Called on 11-14-18. I think they will still use us. She said they used Hexagon because a tracker had a problem.
| Accurus-- | Followed up and scheduled |  |  | Total | FOOT | Failed
| Raptor-- | 2 yr interval. Not needed |  |  | 68 | 13 | 5
[5 empty] | Rate: | 19.12% | 7.353%
Count | Date | Model | Company | SN | FOOT? | Pass? | Ambient temp | Notes | Row Labels
1 | 2017-03-02 | Leica AT901MR | Galaxy Tool | 3654 | NO | Pass | 70.6 |  | Fail
2 | 2017-04-04 | Leica AT901B | Galaxy Tool | 4403 | No | Pass | 71.3 |  | 60066
3 | 2017-04-24 | API LTS-3000 | D-J | LTS34970207 | Yes | Fail | 76.3 |  | Accurus
4 | 2017-05-02 | API LTS-3000 | D-J | LTS36761008 | No | Pass | 76.3 |  | 70215
5 | 2017-06-19 | Leica AT960MR | Dynamic NC | 750064 | Yes | Pass | 71 |  | AIT Technologies
6 | 2017-08-01 | API Radian R50 | Accurus | 60066 | yes | Fail | 71.6 | Followed up. Scheduled. | 750419
7 | 2017-08-08 | Faro Xi | GKN | X02000702066 | No | Pass | 68.4 |  | Raptor
8 | 2017-10-12 | Leica AT960MR | Orizon | 750787 | No | Pass | 68 |  | ?
9 | 2017-10-24 | Faro Ion | Galaxy Tool | Y02001606114 | No | Pass | 68 |  | FMI
10 | 2017-10-30 | API Omnitrac 2 | AIT Technologies | 70215 | Yes | Fail | 77 |  | LTS34970207
11 | 2017-11-01 | Faro Ion | AAT Aero | Y01009003029 | Yes | Pass | 65 | Has Faro Warranty, declined cal in 2018, follow up 2019 | D-J
12 | 2017-11-08 | Leica AT901B | Galaxy Tool | 4176 | No | Pass | 70 |  | Pass
13 | 2017-11-25 | Faro Vantage | Radius Shelbyville | V01001304482 | Yes | Pass | 68.5 |  | 3654
14 | 2017-12-07 | Leica AT930 | Raptor | 750419 | Yes | Fail | 69 | 11-14-18Followed up. On a 2 year interval. Follow up 2019 | Galaxy Tool
15 | 2018-03-05 | Leica AT901 MR | Galaxy Tool | 3654 | NO | Pass | 66 |  | 4176
16 | 2018-04-03 | Leica AT901B | Galaxy Tool | 4403 | No | Pass | 66.7 |  | Galaxy Tool
18 | 2018-05-03 | Leica AT960MR | Galaxy Tool | 751074 | Yes | Pass | 72.5 |  | 4403
17 | 2018-05-03 | Leica AT960MR | Galaxy Tool | 751078 | No | Pass | 68.7 |  | Galaxy Tool
19 | 2018-05-17 | Faro Ion | FMI | Y01001203972 | Yes | Fail | 69.8 |  | 750064
20 | 2018-08-10 | Faro Xi | GKN | X02000702066 | No | Pass | 70.4 | Discovered and corrected CTE Issue | Dynamic NC
21 | 2018-08-21 | Leica AT930 | Vector Tooling Technologies | 750245 | Yes | Pass | 70.3 |  | 750245
22 | 2018-09-05 | Leica AT960MR | Aero-Tech Engineering | 751189 | No | Pass | 74.8 |  | Vector Tooling Technologies
23 | 2018-10-11 | Leica AT960MR | Orizon | 750787 | No | Pass | 69 |  | 750787
24 | 2018-10-23 | Leica AT901B | Galaxy Tool | 4176 | No | Pass | 68.9 |  | Orizon
25 | 2018-10-23 | Faro Ion | Galaxy Tool | Y02001606114 | No | Pass | 71.2 |  | 751074
26 | 2018-11-08 | Faro Vantage | Radius Shelbyville | V01001304482 | No | Pass | 66.2 |  | Galaxy Tool
27 | 2018-11-15 | API Radian R50 | Accurus | 60066 |  |  |  |  | 751078
28 | 2019-01-18 | Leica AT901B | Capps Manufacturing | 3268 | NO | Pass | 60.2 |  | Galaxy Tool
29 | 2019-03-11 | Leica LTD 840 | Bombardier Learjet | 2077 | NO | Pass | 68.5 |  | 751189
30 | 2019-04-01 | Leica AT901B | Galaxy Tool | 4403 | NO | Pass |  | Replaced with Faro Tracker - Due March 3 2021 | Aero-Tech Engineering
31 | 2019-04-01 | Leica AT960MR | Galaxy Tool | 751074 | NO | Pass |  | Repaired - Due 3-27-21 | LTS36761008
32 | 2019-04-16 | Leica AT960MR | Tect Power | 750937 | No | Pass |  |  | D-J
33 | 2019-05-02 | Leica AT901LR | D-J | 3212 | No | Pass |  |  | V01001304482
34 | 2019-05-10 | Leica AT901LR | D-J | 3900 | No | Pass |  |  | Triumph Shelbyville
35 | 2019-06-25 | API Omnitrac 2 | D-J | 70457 | No | Pass |  |  | X02000702066
36 | 2019-06-25 | API Omnitrac 2 | D-J | 70459 | No | Pass |  |  | GKN
37 | 2019-06-25 | API T3 | D-J | 3676 | No | Pass |  |  | Y01009003029
38 | 2019-06-28 | Faro Ion | FMI | Y01001203972 | Yes | Pass |  |  | AAT Aero
39 | 2019-07-24 | Faro Vantage | Bombardier Learjet | V01001304533 | Yes | Pass |  |  | Y02001606114
40 | 2019-08-20 | Leica AT960MR | Vector Tooling Technologies | 750245 | No | Pass |  |  | Galaxy Tool
41 | 2019-08-22 | Faro Xi | GKN | X02000702066 | No | Pass |  |  | Grand Total
42 | 2019-09-03 | Leica AT960MR | Aero-Tech Engineering | 751189 | Yes | Pass
43 | 2019-10-02 | Leica LTD 709 | Radius Hot Springs | 2028 | No | Pass
44 | 2019-10-10 | Leica AT960MR | Orizon | 750787 | No | Pass
45 | 2019-10-25 | Leica AT901B | Galaxy Tool | 4176 | No | Pass |  | Replaced with Faro Tracker - Due March 3 2021
46 | 2019-10-25 | Faro Ion | Galaxy Tool | Y02001606114 | No | Pass |  | Recall SN 751078 in July 2020
47 | 2019-11-06 | Faro Vantage | Radius Shelbyville | V01001304482 | No | Pass
48 | 2019-12-05 | Leica AT930 | Raptor | 750419 | No | Pass
49 | 2020-01-06 | Leica AT901B | Capps Manufacturing | 3268 | No | Pass
50 | 2020-03-10 | Leica LTD 840 | Bombardier Learjet | 2077 | No | Pass
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65
66
67
68

## Software Notes
PC Dmis
To set up point to point measurement measurements
Edit Menu- Preferences- Setup - and select Point only measurements (or something like that)
(you can press F5 to get that menu)
Make sure the 1.5 SMR is selected as the probe. Should be selectable from a top of the screen drop down above the other menus.
Ctrl + H | takes a point
after two points select red button with <---> symbol (distance measurement). Select the two points from the resulting menu and select 3D distance

## tracker specs
Volumetric
|  |  |  | Metric Tolerances |  |  |  | Imperial tolerances
Tolerance
| Notes |  | Tolerance as a formula | 1.524 | 2.7432 | 6.096 | 12.1919 | 5 | 9 | 20 | 40 | 16
Model
1 | For "Horizontal Scale bar measurement" only per B89 4.19 | Faro Ion |  | 0.044 | 0.064 | 0.099 | 0.17 | 0.0017 | 0.0025 | 0.0039 | 0.0067 | 0.0025
2 | For "Horizontal Scale bar measurement" only per B89 4.19 | Faro Xi |  | 0.032 | 0.046 | 0.068 | 0.11 | 0.0013 | 0.0018 | 0.0027 | 0.0043 | 0.0018
3 | Horizontal Scale Bar Measurements Per B89 | Faro Vantage |  | 0.044 | 0.064 | 0.099 | 0.17 | 0.0017 | 0.0025 | 0.0039 | 0.0067 | 0.0025
4 | Scale bar measurement per B89.4.19 | API Radian |  | 0.028 | 0.049 | 0.085 | 0.156 | 0.0011 | 0.0019 | 0.0033 | 0.0061 | 0.0019
5 | Scale bar measurement per B89.4.19 | API Omnitrac 2 |  | 0.035 | 0.057 | 0.092 | 0.163 | 0.0014 | 0.0022 | 0.0036 | 0.0064 | 0.0022
6 | cobbled together component tolerances | API T3 | 3.5µm / m + 5 ppm + 15µm | 0.047074 | 0.063533 | 0.1088 | 0.19109 | 0.0019 | 0.0025 | 0.0043 | 0.0075 | 0.0025
7 | For "Horizontal Scale bar measurement" only per B89 4.19 | Leica AT960 MR |  | 0.036 | 0.063 | 0.106 | 0.191 | 0.0014 | 0.0025 | 0.0042 | 0.0075 | 0.0025
8 | For "Horizontal Scale bar measurement" only per B89 4.20 | Leica AT901 MR |  | 0.06 | 0.06 | 0.06 | 0.28 | 0.0024 | 0.0024 | 0.0024 | 0.011 | 0.0024
9 | For "Horizontal Scale bar measurement" only per B89 4.21 | Leica AT901B | 15µm + 5µm/m | 0.06 | 0.06 | 0.06 | 0.28 | 0.0024 | 0.0024 | 0.0024 | 0.011 | 0.0024
10 | For "Horizontal Scale bar measurement" only per B89 4.22 | Leica AT930 |  | 0.036 | 0.063 | 0.106 | 0.191 | 0.0014 | 0.0025 | 0.0042 | 0.0075 | 0.0025
11 | For "Horizontal Scale bar measurement" only per B89 4.22 | Leica AT403 |  | 0.046 | 0.06 | 0.09 | 0.133 | 0.0018 | 0.0024 | 0.0035 | 0.0052
12 | [7 empty] | 0 | 0 | 0 | 0
13 | [7 empty] | 0 | 0 | 0 | 0
14 | [7 empty] | 0 | 0 | 0 | 0
15 | [7 empty] | 0 | 0 | 0 | 0
16 | [7 empty] | 0 | 0 | 0 | 0
17 | [7 empty] | 0 | 0 | 0 | 0
18
19 | 80% of the largest manufacturer's tolerance for each distance | Industry Tolerance | [5 empty] | 0.00192 | 0.002 | 0.00344 | 0.0088 | 0.002
Make distances variable on Testing worksheet
Add 5th column to table
Make it so whatever distance you pick falls into the actual spec table for each instrument
FIX 16 ft distance
Incorporate backsight tests into volumetric process

## B89.4.19 Testing
[5 empty] | 0.0082 | mm |  | Reference Length Value: | 90.5402 | [5 empty] | Chris | 2.2997 | meters
Length Measurement System Tests |  |  |  |  | 0.00032283 | inches |  | MPE (tolerance) | 0.002 |  | Ref. Length CTE | 0.7 | µin/in/°F |  | Result | 90.5401 | inches
7e-07
| Artifact | Artifact Distance | Rotational Position | [7 empty] | Ref. Length |  |  |  | JV Result
Test Position | Orientation | (Approximate) | of tracker head | Ref. Length | Max Dev. | Tolerance | Result (P/F) | Result 1 | Result2 | Result3 | Temperature | CTE Correction applied |  |  | 2.2997 | meters
1 | Horizontal (Table 2) | 0.23m / 9in | Any (0° default) | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0 |  |  | 90.5402 | inches
2 | Horizontal (Table 2) | 2.8m / 9ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
3 | Horizontal (Table 2) | 2.8m / 9ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
4 | Horizontal (Table 2) | 2.8m / 9ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
5 | Horizontal (Table 2) | 2.8m / 9ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
6 | Vertical (Table 3) | 2.8m / 9ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
7 | Vertical (Table 3) | 2.8m / 9ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
8 | Vertical (Table 3) | 2.8m / 9ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
9 | Vertical (Table 3) | 2.8m / 9ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
10 | Right Diagonal 45° (Table 4) | 2.8m / 9ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
11 | Right Diagonal 45° (Table 4) | 2.8m / 9ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
12 | Right Diagonal 45° (Table 4) | 2.8m / 9ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
13 | Right Diagonal 45° (Table 4) | 2.8m / 9ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
14 | Left Diagonal 45°(Table 5) | 2.8m / 9ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
15 | Left Diagonal 45°(Table 5) | 2.8m / 9ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
16 | Left Diagonal 45°(Table 5) | 2.8m / 9ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
17 | Left Diagonal 45°(Table 5) | 2.8m / 9ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
18 | Horizontal (Table 2) | 6.2m / 20ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
19 | Horizontal (Table 2) | 6.2m / 20ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
20 | Horizontal (Table 2) | 6.2m / 20ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
21 | Horizontal (Table 2) | 6.2m / 20ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
22 | Vertical (Table 3) | 6.2m / 20ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
23 | Vertical (Table 3) | 6.2m / 20ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
24 | Vertical (Table 3) | 6.2m / 20ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
25 | Vertical (Table 3) | 6.2m / 20ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
26 | Right Diagonal 45° (Table 4) | 6.2m / 20ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
27 | Right Diagonal 45° (Table 4) | 6.2m / 20ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
28 | Right Diagonal 45° (Table 4) | 6.2m / 20ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
29 | Right Diagonal 45° (Table 4) | 6.2m / 20ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
30 | Left Diagonal 45°(Table 5) | 6.2m / 20ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
31 | Left Diagonal 45°(Table 5) | 6.2m / 20ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
32 | Left Diagonal 45°(Table 5) | 6.2m / 20ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
33 | Left Diagonal 45°(Table 5) | 6.2m / 20ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
User Selected( Pos. 1) | Vertical, lower nest at tracker height | Close | 0° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
User Selected( Pos. 2) | Compound 45 ° | 5m / 16ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail |  |  |  |  | 0
TWO FACE SYSTEMTESTS
|  | Artifact Distance | Rotational Position | [5 empty] | Repetition 1 |  |  |  | Repetition 2 |  |  |  | Repitition 3
Test Position | Nest Position / Mode | (Approximate) | of tracker head |  |  |  |  | X | Y | Z |  | X | Y | Z |  | X | Y | Z
1 | Low / Frontsight | * | 0° | 90.5402 | -90.5402 | 0.002 | Fail
2 | Low / Backsight | * | 0° | 90.5402 | -90.5402 | 0.002 | Fail
3 | Mid / Frontsight | * | 0° | 90.5402 | -90.5402 | 0.002 | Fail
4 | Mid / Backsight | * | 0° | 90.5402 | -90.5402 | 0.002 | Fail
5 | High / Frontsight | * | 0° | 90.5402 | -90.5402 | 0.002 | Fail
6 | High / Backsight | * | 0° | 90.5402 | -90.5402 | 0.002 | Fail
7 | Low / Frontsight | * | 90° | 90.5402 | -90.5402 | 0.002 | Fail
8 | Low / Backsight | * | 90° | 90.5402 | -90.5402 | 0.002 | Fail
9 | Mid / Frontsight | * | 90° | 90.5402 | -90.5402 | 0.002 | Fail
10 | Mid / Backsight | * | 90° | 90.5402 | -90.5402 | 0.002 | Fail
11 | High / Frontsight | * | 90° | 90.5402 | -90.5402 | 0.002 | Fail
12 | High / Backsight | * | 90° | 90.5402 | -90.5402 | 0.002 | Fail
13 | Low / Frontsight | * | 180° | 90.5402 | -90.5402 | 0.002 | Fail
14 | Low / Backsight | * | 180° | 90.5402 | -90.5402 | 0.002 | Fail
15 | Mid / Frontsight | * | 180° | 90.5402 | -90.5402 | 0.002 | Fail
16 | Mid / Backsight | * | 180° | 90.5402 | -90.5402 | 0.002 | Fail
17 | High / Frontsight | * | 180° | 90.5402 | -90.5402 | 0.002 | Fail
18 | High / Backsight | * | 180° | 90.5402 | -90.5402 | 0.002 | Fail
19 | Low / Frontsight | * | 270° | 90.5402 | -90.5402 | 0.002 | Fail
20 | Low / Backsight | * | 270° | 90.5402 | -90.5402 | 0.002 | Fail
21 | Mid / Frontsight | * | 270° | 90.5402 | -90.5402 | 0.002 | Fail
22 | Mid / Backsight | * | 270° | 90.5402 | -90.5402 | 0.002 | Fail
23 | High / Frontsight | * | 270° | 90.5402 | -90.5402 | 0.002 | Fail
24 | High / Backsight | * | 270° | 90.5402 | -90.5402 | 0.002 | Fail
25 | Low / Frontsight | 3m / 10ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail
26 | Low / Backsight | 3m / 10ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail
27 | Mid / Frontsight | 3m / 10ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail
28 | Mid / Backsight | 3m / 10ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail
29 | High / Frontsight | 3m / 10ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail
30 | High / Backsight | 3m / 10ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail
31 | Low / Frontsight | 3m / 10ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail
32 | Low / Backsight | 3m / 10ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail
33 | Mid / Frontsight | 3m / 10ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail
34 | Mid / Backsight | 3m / 10ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail
35 | High / Frontsight | 3m / 10ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail
36 | High / Backsight | 3m / 10ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail
37 | Low / Frontsight | 3m / 10ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail
38 | Low / Backsight | 3m / 10ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail
39 | Mid / Frontsight | 3m / 10ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail
40 | Mid / Backsight | 3m / 10ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail
41 | High / Frontsight | 3m / 10ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail
42 | High / Backsight | 3m / 10ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail
43 | Low / Frontsight | 3m / 10ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail
44 | Low / Backsight | 3m / 10ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail
45 | Mid / Frontsight | 3m / 10ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail
46 | Mid / Backsight | 3m / 10ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail
47 | High / Frontsight | 3m / 10ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail
48 | High / Backsight | 3m / 10ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail
49 | Low / Frontsight | 6m / 20ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail
50 | Low / Backsight | 6m / 20ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail
51 | Mid / Frontsight | 6m / 20ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail
52 | Mid / Backsight | 6m / 20ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail
53 | High / Frontsight | 6m / 20ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail
54 | High / Backsight | 6m / 20ft | 0° | 90.5402 | -90.5402 | 0.002 | Fail
55 | Low / Frontsight | 6m / 20ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail
56 | Low / Backsight | 6m / 20ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail
57 | Mid / Frontsight | 6m / 20ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail
58 | Mid / Backsight | 6m / 20ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail
59 | High / Frontsight | 6m / 20ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail
60 | High / Backsight | 6m / 20ft | 90° | 90.5402 | -90.5402 | 0.002 | Fail
61 | Low / Frontsight | 6m / 20ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail
62 | Low / Backsight | 6m / 20ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail
63 | Mid / Frontsight | 6m / 20ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail
64 | Mid / Backsight | 6m / 20ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail
65 | High / Frontsight | 6m / 20ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail
66 | High / Backsight | 6m / 20ft | 180° | 90.5402 | -90.5402 | 0.002 | Fail
67 | Low / Frontsight | 6m / 20ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail
68 | Low / Backsight | 6m / 20ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail
69 | Mid / Frontsight | 6m / 20ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail
70 | Mid / Backsight | 6m / 20ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail
71 | High / Frontsight | 6m / 20ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail
72 | High / Backsight | 6m / 20ft | 270° | 90.5402 | -90.5402 | 0.002 | Fail
*As Close as possible to the tracker to maximize vertical encoder use between the end nests
IFM RANGING TEST | [8 empty] | Reference length CTE |  | µin/in/°F
| Artifact | Artifact Distance | Rotational Position | [7 empty] | Ref. Length
Test Position | Orientation | (Approximate) | of tracker head |  |  |  |  | Result 1 | Result2 | Result3 | Temperature
1 | Pure Radial | 0.575m / 22in | 0° | 22.625 | -22.625 | 0.0005 | Fail |  |  |  |  | 0
2 | Pure Radial | 1.15m / 44 in | 0° | 45 | -45 | 0.0005 | Fail |  |  |  |  | 0
3 | Pure Radial | 1.725m / 67 in | 0° | 67 | -67 | 0.0005 | Fail |  |  |  |  | 0
4 | Pure Radial | 2.3m / 90 in | 0° | 90.5 | -90.5 | 0.0005 | Fail |  |  |  |  | 0
ADM RANGING TEST
| Artifact | Artifact Distance | Rotational Position | [7 empty] | Ref. Length
Test Position | Orientation | (Approximate) | of tracker head |  |  |  |  | Result 1 | Result2 | Result3 | Temperature
1 | Pure Radial | 18%R | 0° |  | 0 | 0.0005 | Pass
2 | Pure Radial | 36%R | 0° |  | 0 | 0.0005 | Pass
3 | Pure Radial | 54%R | 0° |  | 0 | 0.0005 | Pass
4 | Pure Radial | 72%R | 0° |  | 0 | 0.0005 | Pass
5 | Pure Radial | 10%R | 0° |  | 0 | 1.0005 | Pass
6 | Pure Radial | 25%R | 0° |  | 0 | 2.0005 | Pass

## Sheet2
| Vertical Deviation | Horizontal
Distance
5 ft, 0° | -0.005 | 0.002
9 ft, 0° | -0.0053 | 0.0005
9 ft, 90° | -0.0047 | 0.0003
9 ft, 180° | -0.0066 | 0.002
9 ft, 270° | -0.0054 | 0.0015
20 ft, 0° | -0.0045 | 0.0018
40 ft, 0° | -0.0057 | 0.0011

## Logo Page
A2LA
1

## Software Validation
Date | Action
2017-06-23 | initial validation performed
2018-09-01 | Full certificate brought into Excel. Tolerance worksheet established and validated
2018-10-12 | Corrected errant references in SMR sheet, applied units on cert page.
2018-11-01 | Added Radial reading to datasheet, also added detailed backsight results